## Slide 1: Staging Dutch magnet summary
2018/7/1, ver. 1
Kei Nakamura

## Slide 2: Revision
2018/7/1
  Staging Dutch is now Moore Dutch 1 (2 after undulator planned)
2014/3/25
  Cam1 was removed for cross calibration. Before and after this operation, camera position was saved in
  Before: 3/21 scan4
  After: 3/25 scan1
  Before data is to confirm the drift since installation (2011). After is to measure the change due to removal and re-installation.

## Slide 3: 7/1/2018 CCD camera 1 (front)
21.5 mm
Beam tube center
10 mm
130mm/555pixel, 0.234 mm/pixel
Field of view (692pix) = 162.1 mm
Image right edge = 7.9 mm
Range: 7.9 – 170 mm
X=634 pix
Data: 2018_0628 scan 11 and 12
Notes: X63 – 618, 130 mm, right: 58 pixels from 21.5mm, 7.9 mm

## Slide 4: 7/1/2018 CCD camera 2 (downstream, side)
X=683 pix
332.4 mm from center
160mm/573pixel, 0.279 mm/pixel
Field of view = 193.2 mm
Range: 141.7 – 334.9 mm
150.2 mm from center of magnet
Data: 2018_0628 scan 11 and 12
Notes: X: 644, 71, 160 mm, right: 9 pixel from 332.4

## Slide 5: 7/1/2018 CCD camera 3 (upstream, side)
X=657 pix
150.2 mm from center of magnet
160mm/573pixel, 0.279 mm/pixel
Field of view = 193.2 mm
Range: -33.2 – 160 mm
Save ROI: top 184, height 130
Data: 2018_0628 scan 11 and 12
Notes: 658-85=573 pixel, 160mm, 0.279

## Slide 6: Chamber is 5.7 mm too deep in the yoke
Designed
Measured
Probe holder
129.2mm
Notes: Measured on 7/10/2018, Jeroen

## Slide 7: Beam line and screen wrt the magnet
Chamber (screen)
Pole
Z[mm]
+x
Notes
Magnet can float with respect to the chamber (magnet center can have offset with respect to the beam line).
The chamber defines the beam line and screen location.
Front screen is at z =354.6 mm, and side screen is at x = 170 mm.
The chamber end-end is 27.925” = 709.3mm
Xm=5.7
Beam line
170 mm
0
354.6
-354.6
Magnet center
240
-240
20.5 (chamber wall)
Yoke wall
129.2 (134.9mm when
beam line = magnet center)
21.5
157.2
332.4
-332.4

## Slide 8: Peak field ~1600 mT
Magnetic field is always monitored/recorded by a Hall probe.
Control for the power supply is …
Data: server/TREXResources/calibration/DutchMag/100722BIScan.xlsx
NOT updated

## Slide 9: Effective field size 144.4 x 504.4 mm^2
Effective boundary
Effective boundary
Magnet edge
r
g
12 cm  48 cm
x
Pole size is 120 mm x 480 mm,
Effective field size is 144.4 mm x 504.4 mm.
Data: server/TREXResources/calibration/DutchMag/040330BXZScan.xlsx

## Slide 10: Archives

## Slide 11: 2 GeV, 7.0256 degree, 41 mm
Field
Coil
Displacement @ outside of the coil
With 1.7 T, it bends 2 GeV beam 7 degree.
Program: server/software/matlab/Kei/programs/elcTrajectory/dutch/dutchTrjV01.m
NOT updated

## Slide 12: Staging beam line design

## Slide 13
Z=0 [mm]
Laser focus
Z=33 [mm]
Nominal cap1 exit
Z=96 [mm]
Nominal cap2 exit
Z=2009 [mm]
Chamber input
I.D. = 15.24 mm
Z=2363 [mm]
Magnet center
Z [mm]
Z=43 [mm]
Plasma mirror

## Slide 14: Front screen momentum-position
3.8645e-9*x^6 -2.5268e-6*x^5 +6.7217e-4*x^4 -0.093486*x^3 +7.2584*x^2 -310.253*x +6656.9
Data: server/TREXResources/calibration/DutchMag/frontScreen.xlsx
Program: server/software/matlab/Kei/programs/elcTrajectory/dutch/dutchTrjV02.m
NOT updated

## Slide 15: Side screen momentum-position
-4.6091e-10*x^4 +166828e-7*x^3 +7.485e-4*x^2 +0.37296*x +66.178
Data: server/TREXResources/calibration/DutchMag/sideScreen.xlsx
NOT updated

## Slide 16: E-beam acceptance
Laser focus
0
~1900
Dutch magnet front
~2800
Beam block
~3100 mm
Dutch magnet end
Aperture @ Dutch magnet exit: I.D. 7/8” = 22.225 mm
Aperture @ beam tube near entrance: I.D. = 15.24 mm
E-beam acceptance: 15.24/2.8 ~ 5.4 mrad (+/- 2.7 mrad)

## Slide 17: 3/21/2014 CCD camera 1 (front), no change was observed.
21.5 mm
Beam tube center
10 mm
120mm/520pixel, 0.231 mm/pixel
Field of view (692pix) = 159.7 mm
Image right edge = 10.7 mm
Range: 10.7 – 170.4 mm
X=645 pix
Notes: Note: picture is old

## Slide 18: 3/21/2014 CCD camera 2 (downstream, side), no change
X=666 pix
332.4 mm from center
Data: Server/5_resources/StagingResources/Staging van der spec calibration/side cam1
180mm/638pixel, 0.282 mm/pixel
Field of view = 195.2 mm
Range: 144.5 – 339.7 mm
Save ROI: top 180, height 130
X=91 pix
170.2 mm from center of magnet

## Slide 19: 3/21/2014 CCD camera 3 (upstream, side), no change
X=668 pix
170.2 mm from center of magnet
Data: Server/5_resources/StagingResources/Staging van der spec calibration/side cam2
180mm/624pixel, 0.288 mm/pixel
Field of view = 199.6 mm
Range: -22.5 – 177.1 mm
Save ROI: top 184, height 130

## Slide 20: 3/25/2014 CCD camera 1 (front), changed, calibration is modified.
21.5 mm
Beam tube center
10 mm
Note: image is old
120mm/513pixel, 0.234 mm/pixel
Field of view (692pix) = 161.9 mm
Image right edge = 3.25 mm
Range: 3.25 – 165.2 mm
X=614 pix

## Slide 21: 3/25/2014 CCD camera 2 (downstream, side), no change
X=666 pix
332.4 mm from center
Data: Server/5_resources/StagingResources/Staging van der spec calibration/side cam1
180mm/638pixel, 0.282 mm/pixel
Field of view = 195.2 mm
Range: 144.5 – 339.7 mm
Save ROI: top 180, height 130
X=91 pix
170.2 mm from center of magnet

## Slide 22: 3/25/2014 CCD camera 3 (upstream, side), no change
X=668 pix
170.2 mm from center of magnet
Data: Server/5_resources/StagingResources/Staging van der spec calibration/side cam2
180mm/624pixel, 0.288 mm/pixel
Field of view = 199.6 mm
Range: -22.5 – 177.1 mm
Save ROI: top 184, height 130